10525 Willows Road NE

Redmond, WA 98052

Phone: 1.425.881.6444

Phone: 1.800.247.5700 (press 2)

Fax: 425.867.6972

pmhelp@data-io.com

### SUBJECT: PP100 Hour Meter Replacement in the field.

| MODELS AFFECTED |  | 2500 |  | 3000 |  | 7000 |  | 7500 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | #### PP100 | X | 9500 |  | 970 |  | T/R |  | ALL |  |

1.0	PURPOSE: To determine if the Hour Meter mounted on the PP100 back panel is to be replaced and if so, replace it with a new one.

- Tools required:

Flat screw driver and nut driver or pliers.

- Equipment Required:

  - New Hour Meter, Data I/O P/N: 384-0003-001

ENM Chicago, P/N T50B1 (230V, 50/60Hz)

###### Connector cable kit, Data I/O P/N: xxx-xxxx-xxx

- Instructions:-

  - Determine if the installed Hour Meter has to be replaced:-

Hour meters with push-button RESET have to be replaced. The reset button is on the left of hour digits and is manufactured by CRAMER P/N 636W-AA 115V, 60Hz – this needs to be replaced by the new type meter.

In USA, hour meters which do not have the reset button need not be replaced.

However, in countries with 50Hz AC supply all Hour Meters need to be replaced. The older ones, rated at 60Hz, are frequency dependant and may not be giving correct hour count.

  - Remove the old Hour Meter

##### Turn-off power to PP100 and remove the power supply coming into the PP100.

##### Disconnect the electrical connector coming into the Hour Meter and remove the Hour Meter from the PP100 Back Panel. Save the three screws, their washers and the front (black) plastic bezel. Discard the metal spacer ring (silver colored ring) – this will not be required.

  - Install the New Hour meter, which does NOT have the Reset Button

Insert the New Hour Meter into the hole of the Back Panel.

Put the black bezel (removed from the old Hour Meter) on top of the New Hour Meter.

Screw down the bezel on the Back Panel, with the New Hour Meter in between the bezel and the Back Panel.

Connect the red wire of the connector to the lug marked + on the New Hour Meter and the Black wire to the lug marked – on the meter.

Re-reconnect the other end of the wire/connector to existing connector.

Power-up the PP100 and verify that the New Hour Meter is working.